## Spek Teknis Sebelum
SPESIFIKASI TEKNIS
Pengadaan Alat-alat Laboratorium Sentral dan Prodi-prodi di lingkungan Unand
NO. | NAMA BARANG | SPESIFIKASI |  |  | Volume | Satuan
15 | AAS (Atomic Absorption Spectrophotometer) | Specifications : |  |  | 1 | Unit
|  | Number of HC lamps | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp,litting one.
|  | Monochromator | : | Aberration-corrected Czerny-Turner mounting
|  | Optics | : | Optical double beam
|  | Wavelength range | : | 185 to 900 nm
|  | Number of grating | : | 1800 lines, Focal lenght 300 mm
|  | Focal length | : | 300 mm
|  | Bandwidth | : | 0.1, 0.2, 0.4, 0.7, 1.0, 2.0 nm (6-step automatic switching)
|  | Detector | : | Photomultiplier tube (PMT)
|  | Lamp. | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp, litting one.
|  | External PC | : | Fully Computer Controlled System
|  | Background Correction Method | : | · BGC-SR (high-speed self-reversal method) · BGC-D2 (D2 lamp method)
FLAME TYPE
|  | Burner Head | : | Titanium 10cm slot (5cm titanium slot for N2O-C2H2 flame available as an option)
|  | Flow rate control | : | Automatic fuel gas flow rate setting (0.1 L/min step)
Gas Control
|  | * Flow rate control |  | Fuel: Automatic search for optimum flow rate (0.1 L/min step). Automatic search of optimum gas flow rate
Support: Manual adjustment with flow meter
|  | * Safety features |  | ● Automatic gas leak check ● Automatic switching between Air-C2H2 and N2O-C2H2 ● Flame monitor ● Prevention of wrong burner head use ● Gas pressure monitor ● Drain tank level monitor ● Automatic flame extinction upon power outage or sudden power interruption ● Automatic flame extinction via flame vibration sensor ● Internal fan stop sensor
Safety Features
|  | Positioning | : | ● Lateral / Vertical Adjustment
|  | Power Requirements | : | 230 or 240 VAC, + - 5 % , 7400 VA, 50/60 Hz
Autosampler
|  | Control | : | RS-232 C Communication control from AA Unit
|  | Maximum reagent/sample position | : | Reagents : 8 positions Samples : 60 position
(random access available with reagents or samples)
|  | Rinse Water Bottle | : | 2 L
|  | Power Requirements | : | Choose from 100, 120,220 or 230 VAC, 50 VA, 50/60 Hz
|  | Function | : | Zero-point detection, auto diagnosis,auto rinse and random access
|  | Sample Volume | : | Sample vials : 15 mL, Reagents Vials : 53 mL
|  | Nozzle Rinse | : | Solvent rinse aspiration method
Software:
Compatible with Windows
10x Hollow Cathode Lamps with Standard Solutions
1x Pressure Regulator Acetylene + Gas Acetylene with Cylinder
1x Pressure Regulator N2O + Gas N2O with Cylinder
1x Oil-less Air Compressor Piston Type with Vent and Ducting
Desktop, PC HP 280 G4 MT, Intel Core i5 (2.8GHz up to 4.0 GHz), Ram 4GB D4-2400 DIMM (2Slot up to 32GB), HDD 1TB 7200RPM, Free DOS / Windows 10 Pro 64, HP LED Monitor 19", HP ProDisplay P19A 19-inch LED Backlit Monitor, 640x480 @ 60 Hz/ 1280x1024 @ 60 Hz720x400 @ 70 Hz/ 800x600 @ 60 Hz/ 1024x768 @ 60 Hz/ 1280x960 @ 60 Hz Resolutions, 5:4 Aspect ratio, VGA Input signal. Epson L3110 All-in-one Printer
Installation Instrument
Manual Book (Hardcopy and Softcopy)
Training Operator
Certificate of Origin
20 | Particle Analyzer | General Specifications |  |  | 1 | Unit
|  | Measurement principle |  | Laser Diffraction Method
|  | Measurement range |  | 17nm(0.017μm) to 2500μm
17nm(0.017μm) to 400μm
30nm(0.030μm) to 280μm
300nm(0.3μm) to 2500μm
Measurement Unit
|  | Light source |  | Red Semiconductor laser (Wavelength 680 nm)
|  | Light detector |  | Detector elements for UV semiconductor laser
Total 84 elements (78 forward, 1 side, 5 back)
|  | System Compliance |  | Class 1 Laser Product, CE compliant
|  | Required power supply |  | 115 or 230 VAC as ordered 100 VA
|  | Dimensions & weight |  | W680mm×D280mm×H430mm, 31kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Sampler
|  | Dispersing bath |  | Capacity: 100～280cm3
|  | Sonicator |  | Frequency about 32 kHz, output about 40 W
|  | Liquid Pump |  | Radial pump, maximum flow rate 2000cm3/min
|  | Liquid Pump Material |  | Stainless (SUS 304, SUS 316), Tetrafluoroethylene (PTFE),
Perfluoroelastmor (FEP) or Kalrez, Thermoflon Pascal (inside)
|  | Liquid Supply Pump |  | Diaphragm pump, maximum flow rate 750cm3/min
|  | Liquid Supply Pump Material |  | Tetrafluoroethylene, polyvinyl dene fluoride
|  | Flow Cell |  | Quartz glass
|  | Required power supply |  | 115 or 230 VAC as ordered, 200 VA
|  | Dimensions & weight |  | W390mm×D520mm×H430mm,18kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Batch Cell
|  | Cell Material |  | Quartz glass
|  | Required Liquid Volume |  | Approx. 12 cm3
|  | Stirrer Mechanism |  | Up-and-down movement of blade
|  | Dimensions & weight |  | W100mm×D120mm×H140mm, 0.8kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
High-concentration Sample Measurement System
|  | Cell Material |  | Borosilicate glass
|  | Required Liquid Volume |  | Approx. 0.15 cm3
|  | Dimensions & weight |  | W20mm×D100mm×H9mm, 0.2kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Cyclone Injection Type Dry Measurement Unit
|  | Sample suction types | Cyclone type / One shot type / Hand shot type
|  | Suction nozzle | Can be selected from 3 types
Sampling Unit Specifications
|  | System | Cyclone type
|  | Communication method | USB (PC control)
|  | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Dimensions & weight | W 240×D 310×H 210 mm, 10 kg
|  | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Pressure Regulator Specifications
|  | Primary pressure | 0.6 to 0.8 MPa
|  | Secondary pressure | 0.05 to 0.5 MPa
|  | Filtration rating | Removing particles of 5 μm or larger
|  | Connection to air source | Tube of 6 mm outside diameter
|  | Communication method | USB (PC control)
|  | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Dimensions & weight | W 130×D 223×H 233 mm, 3 kg
|  | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)

## Spek Teknis Sesudah
SPESIFIKASI TEKNIS
PENGADAAN ALAT LABORATORIUM UNIVERSITAS ANDALAS TAHUN 2019
NO. | NAMA BARANG | SPESIFIKASI |  |  | QTY
15 | AAS (Atomic Absorption Spectrophotometer) | Specifications : |  |  | 1 | Unit
Specifications :
|  | Number of HC lamps | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp,litting one.
|  | Monochromator | : | Aberration-corrected Czerny-Turner mounting
|  | Optics | : | Optical double beam
|  | Wavelength range | : | 185 to 900 nm
|  | Number of grating | : | 1800 lines, Focal lenght 300 mm
|  | Focal length | : | 300 mm
|  | Bandwidth | : | 0.1, 0.2, 0.4, 0.7, 1.0, 2.0 nm (6-step automatic switching)
|  | Detector | : | Photomultiplier tube (PMT)
|  | Lamp. | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp, litting one.
|  | External PC | : | Fully Computer Controlled System
|  | Background Correction Method | : | · BGC-SR (high-speed self-reversal method) · BGC-D2 (D2 lamp method)
FLAME TYPE
|  | Burner Head | : | Titanium 10cm slot (5cm titanium slot for N2O-C2H2 flame available as an option)
|  | Flow rate control | : | Automatic fuel gas flow rate setting (0.1 L/min step)
Gas Control
|  | * Flow rate control |  | Fuel: Automatic search for optimum flow rate (0.1 L/min step). Automatic search of optimum gas flow rate
Support: Manual adjustment with flow meter
|  | * Safety features |  | ● Automatic gas leak check ● Automatic switching between Air-C2H2 and N2O-C2H2 ● Flame monitor ● Prevention of wrong burner head use ● Gas pressure monitor ● Drain tank level monitor ● Automatic flame extinction upon power outage or sudden power interruption ● Automatic flame extinction via flame vibration sensor ● Internal fan stop sensor
|  | Safety Features |  | ● Cooling water flow rate monitor ● Gas Pressure Monitor ● Overcurrent protection unit ( double check by circuit protector and optical sensor ) ● Furnace block cooling check
|  | Positioning | : | ● Lateral / Vertical Adjustment
|  | Power Requirements | : | 230 or 240 VAC, + - 5 % , 7400 VA, 50/60 Hz
Software:
Compatible with Windows
10x Hollow Cathode Lamps with Standard Solutions
1x Pressure Regulator Acetylene + Gas Acetylene with Cylinder
1x Pressure Regulator N2O + Gas N2O with Cylinder
1x Oil-less Air Compressor Piston Type with Vent and Ducting
Desktop, PC HP 280 G4 MT, Intel Core i5 (2.8GHz up to 4.0 GHz), Ram 4GB D4-2400 DIMM (2Slot up to 32GB), HDD 1TB 7200RPM, Free DOS / Windows 10 Pro 64, HP LED Monitor 19", HP ProDisplay P19A 19-inch LED Backlit Monitor, 640x480 @ 60 Hz/ 1280x1024 @ 60 Hz720x400 @ 70 Hz/ 800x600 @ 60 Hz/ 1024x768 @ 60 Hz/ 1280x960 @ 60 Hz Resolutions, 5:4 Aspect ratio, VGA Input signal. Epson L3110 All-in-one Printer
Installation Instrument
Manual Book (Hardcopy and Softcopy)
Training Operator
Certificate of Origin
20 | Particle Analyzer | General Specifications |  |  | 1 | Unit
General Specifications
|  | Measurement principle |  | Laser Diffraction Method
|  | Measurement range |  | 17nm(0.017μm) to 2500μm
17nm(0.017μm) to 400μm
30nm(0.030μm) to 280μm
300nm(0.3μm) to 2500μm
Measurement Unit
|  | Light source |  | Red Semiconductor laser (Wavelength 680 nm)
|  | Light detector |  | Detector elements for UV semiconductor laser
Total 84 elements (78 forward, 1 side, 5 back)
|  | System Compliance |  | Class 1 Laser Product, CE compliant
|  | Required power supply |  | 115 or 230 VAC as ordered 100 VA
|  | Dimensions & weight |  | W680mm×D280mm×H430mm, 31kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
High-concentration Sample Measurement System
|  | Cell Material |  | Borosilicate glass
|  | Required Liquid Volume |  | Approx. 0.15 cm3
|  | Dimensions & weight |  | W20mm×D100mm×H9mm, 0.2kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Cyclone Injection Type Dry Measurement Unit
|  | Sample suction types | Cyclone type / One shot type / Hand shot type
|  | Suction nozzle | Can be selected from 3 types
Sampling Unit Specifications
|  | System | Cyclone type
|  | Communication method | USB (PC control)
|  | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Dimensions & weight | W 240×D 310×H 210 mm, 10 kg
|  | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Pressure Regulator Specifications
|  | Primary pressure | 0.6 to 0.8 MPa
|  | Secondary pressure | 0.05 to 0.5 MPa
|  | Filtration rating | Removing particles of 5 μm or larger
|  | Connection to air source | Tube of 6 mm outside diameter
|  | Communication method | USB (PC control)
|  | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Dimensions & weight | W 130×D 223×H 233 mm, 3 kg
|  | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)

## GABUNGAN
SPESIFIKASI TEKNIS
Pengadaan Alat-alat Laboratorium Sentral dan Prodi-prodi di lingkungan Unand
NO. | NAMA BARANG | SPESIFIKASI SEBELUM ADENDUM |  |  | SPESIFIKASI SESUDAH ADENDUM |  |  | Volume | Satuan
15 | AAS (Atomic Absorption Spectrophotometer) | Specifications : |  |  | Specifications : |  |  | 1 | Unit
|  | Number of HC lamps | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp,litting one. | Specifications :
|  | Monochromator | : | Aberration-corrected Czerny-Turner mounting | Number of HC lamps | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp,litting one.
|  | Optics | : | Optical double beam | Monochromator | : | Aberration-corrected Czerny-Turner mounting
|  | Wavelength range | : | 185 to 900 nm | Optics | : | Optical double beam
|  | Number of grating | : | 1800 lines, Focal lenght 300 mm | Wavelength range | : | 185 to 900 nm
|  | Focal length | : | 300 mm | Number of grating | : | 1800 lines, Focal lenght 300 mm
|  | Bandwidth | : | 0.1, 0.2, 0.4, 0.7, 1.0, 2.0 nm (6-step automatic switching) | Focal length | : | 300 mm
|  | Detector | : | Photomultiplier tube (PMT) | Bandwidth | : | 0.1, 0.2, 0.4, 0.7, 1.0, 2.0 nm (6-step automatic switching)
|  | Lamp. | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp, litting one. | Detector | : | Photomultiplier tube (PMT)
|  | External PC | : | Fully Computer Controlled System | Lamp. | : | 8-lamp turret, 2 lamps simultaneously lit (1 for measurement, 1 warming up for next measurement) 2 - sockets for High Performance booted Lamp, litting one.
|  | Background Correction Method | : | · BGC-SR (high-speed self-reversal method) · BGC-D2 (D2 lamp method) | External PC | : | Fully Computer Controlled System
|  | FLAME TYPE |  |  | Background Correction Method | : | · BGC-SR (high-speed self-reversal method) · BGC-D2 (D2 lamp method)
|  | Burner Head | : | Titanium 10cm slot (5cm titanium slot for N2O-C2H2 flame available as an option) | FLAME TYPE
|  | Flow rate control | : | Automatic fuel gas flow rate setting (0.1 L/min step) | Burner Head | : | Titanium 10cm slot (5cm titanium slot for N2O-C2H2 flame available as an option)
|  | Gas Control |  |  | Flow rate control | : | Automatic fuel gas flow rate setting (0.1 L/min step)
|  | * Flow rate control |  | Fuel: Automatic search for optimum flow rate (0.1 L/min step). Automatic search of optimum gas flow rate | Gas Control
|  |  |  | Support: Manual adjustment with flow meter | * Flow rate control |  | Fuel: Automatic search for optimum flow rate (0.1 L/min step). Automatic search of optimum gas flow rate
|  | * Safety features |  | ● Automatic gas leak check ● Automatic switching between Air-C2H2 and N2O-C2H2 ● Flame monitor ● Prevention of wrong burner head use ● Gas pressure monitor ● Drain tank level monitor ● Automatic flame extinction upon power outage or sudden power interruption ● Automatic flame extinction via flame vibration sensor ● Internal fan stop sensor |  |  | Support: Manual adjustment with flow meter
|  | Safety Features |  |  | * Safety features |  | ● Automatic gas leak check ● Automatic switching between Air-C2H2 and N2O-C2H2 ● Flame monitor ● Prevention of wrong burner head use ● Gas pressure monitor ● Drain tank level monitor ● Automatic flame extinction upon power outage or sudden power interruption ● Automatic flame extinction via flame vibration sensor ● Internal fan stop sensor
|  | Positioning | : | ● Lateral / Vertical Adjustment | Safety Features |  | ● Cooling water flow rate monitor ● Gas Pressure Monitor ● Overcurrent protection unit ( double check by circuit protector and optical sensor ) ● Furnace block cooling check
|  | Power Requirements | : | 230 or 240 VAC, + - 5 % , 7400 VA, 50/60 Hz | Positioning | : | ● Lateral / Vertical Adjustment
[5 empty] | Power Requirements | : | 230 or 240 VAC, + - 5 % , 7400 VA, 50/60 Hz
Autosampler
|  | Control | : | RS-232 C Communication control from AA Unit | Software:
|  | Maximum reagent/sample position | : | Reagents : 8 positions Samples : 60 position | Compatible with Windows
|  |  |  | (random access available with reagents or samples) | 10x Hollow Cathode Lamps with Standard Solutions
|  | Rinse Water Bottle | : | 2 L | 1x Pressure Regulator Acetylene + Gas Acetylene with Cylinder
|  | Power Requirements | : | Choose from 100, 120,220 or 230 VAC, 50 VA, 50/60 Hz | 1x Pressure Regulator N2O + Gas N2O with Cylinder
|  | Function | : | Zero-point detection, auto diagnosis,auto rinse and random access | 1x Oil-less Air Compressor Piston Type with Vent and Ducting
|  | Sample Volume | : | Sample vials : 15 mL, Reagents Vials : 53 mL | Desktop, PC HP 280 G4 MT, Intel Core i5 (2.8GHz up to 4.0 GHz), Ram 4GB D4-2400 DIMM (2Slot up to 32GB), HDD 1TB 7200RPM, Free DOS / Windows 10 Pro 64, HP LED Monitor 19", HP ProDisplay P19A 19-inch LED Backlit Monitor, 640x480 @ 60 Hz/ 1280x1024 @ 60 Hz720x400 @ 70 Hz/ 800x600 @ 60 Hz/ 1024x768 @ 60 Hz/ 1280x960 @ 60 Hz Resolutions, 5:4 Aspect ratio, VGA Input signal. Epson L3110 All-in-one Printer
|  | Nozzle Rinse | : | Solvent rinse aspiration method | Installation Instrument
Manual Book (Hardcopy and Softcopy)
|  | Software: |  |  | Training Operator
|  | Compatible with Windows |  |  | Certificate of Origin
10x Hollow Cathode Lamps with Standard Solutions
1x Pressure Regulator Acetylene + Gas Acetylene with Cylinder
1x Pressure Regulator N2O + Gas N2O with Cylinder
1x Oil-less Air Compressor Piston Type with Vent and Ducting
Desktop, PC HP 280 G4 MT, Intel Core i5 (2.8GHz up to 4.0 GHz), Ram 4GB D4-2400 DIMM (2Slot up to 32GB), HDD 1TB 7200RPM, Free DOS / Windows 10 Pro 64, HP LED Monitor 19", HP ProDisplay P19A 19-inch LED Backlit Monitor, 640x480 @ 60 Hz/ 1280x1024 @ 60 Hz720x400 @ 70 Hz/ 800x600 @ 60 Hz/ 1024x768 @ 60 Hz/ 1280x960 @ 60 Hz Resolutions, 5:4 Aspect ratio, VGA Input signal. Epson L3110 All-in-one Printer
Installation Instrument
Manual Book (Hardcopy and Softcopy)
Training Operator
Certificate of Origin
20 | Particle Analyzer | General Specifications |  |  | General Specifications |  |  | 1 | Unit
|  | Measurement principle |  | Laser Diffraction Method | Measurement principle |  | Laser Diffraction Method
|  | Measurement range |  | 17nm(0.017μm) to 2500μm | Measurement range |  | 17nm(0.017μm) to 2500μm
|  |  |  | 17nm(0.017μm) to 400μm |  |  | 17nm(0.017μm) to 400μm
|  |  |  | 30nm(0.030μm) to 280μm |  |  | 30nm(0.030μm) to 280μm
|  |  |  | 300nm(0.3μm) to 2500μm |  |  | 300nm(0.3μm) to 2500μm
|  | Measurement Unit |  |  | Measurement Unit
|  | Light source |  | Red Semiconductor laser (Wavelength 680 nm) | Light source |  | Red Semiconductor laser (Wavelength 680 nm)
|  | Light detector |  | Detector elements for UV semiconductor laser | Light detector |  | Detector elements for UV semiconductor laser
|  |  |  | Total 84 elements (78 forward, 1 side, 5 back) |  |  | Total 84 elements (78 forward, 1 side, 5 back)
|  | System Compliance |  | Class 1 Laser Product, CE compliant | System Compliance |  | Class 1 Laser Product, CE compliant
|  | Required power supply |  | 115 or 230 VAC as ordered 100 VA | Required power supply |  | 115 or 230 VAC as ordered 100 VA
|  | Dimensions & weight |  | W680mm×D280mm×H430mm, 31kg | Dimensions & weight |  | W680mm×D280mm×H430mm, 31kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation) | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Sampler
|  | Dispersing bath |  | Capacity: 100～280cm3 | High-concentration Sample Measurement System
|  | Sonicator |  | Frequency about 32 kHz, output about 40 W | Cell Material |  | Borosilicate glass
|  | Liquid Pump |  | Radial pump, maximum flow rate 2000cm3/min | Required Liquid Volume |  | Approx. 0.15 cm3
|  | Liquid Pump Material |  | Stainless (SUS 304, SUS 316), Tetrafluoroethylene (PTFE), | Dimensions & weight |  | W20mm×D100mm×H9mm, 0.2kg
|  |  |  | Perfluoroelastmor (FEP) or Kalrez, Thermoflon Pascal (inside) | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
|  | Liquid Supply Pump |  | Diaphragm pump, maximum flow rate 750cm3/min
|  | Liquid Supply Pump Material |  | Tetrafluoroethylene, polyvinyl dene fluoride | Cyclone Injection Type Dry Measurement Unit
|  | Flow Cell |  | Quartz glass | Sample suction types | Cyclone type / One shot type / Hand shot type
|  | Required power supply |  | 115 or 230 VAC as ordered, 200 VA | Suction nozzle | Can be selected from 3 types
|  | Dimensions & weight |  | W390mm×D520mm×H430mm,18kg | Sampling Unit Specifications
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation) | System | Cyclone type
[5 empty] | Communication method | USB (PC control)
|  | Batch Cell |  |  | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Cell Material |  | Quartz glass | Dimensions & weight | W 240×D 310×H 210 mm, 10 kg
|  | Required Liquid Volume |  | Approx. 12 cm3 | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
|  | Stirrer Mechanism |  | Up-and-down movement of blade | Pressure Regulator Specifications
|  | Dimensions & weight |  | W100mm×D120mm×H140mm, 0.8kg | Primary pressure | 0.6 to 0.8 MPa
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation) | Secondary pressure | 0.05 to 0.5 MPa
[5 empty] | Filtration rating | Removing particles of 5 μm or larger
|  | High-concentration Sample Measurement System |  |  | Connection to air source | Tube of 6 mm outside diameter
|  | Cell Material |  | Borosilicate glass | Communication method | USB (PC control)
|  | Required Liquid Volume |  | Approx. 0.15 cm3 | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Dimensions & weight |  | W20mm×D100mm×H9mm, 0.2kg | Dimensions & weight | W 130×D 223×H 233 mm, 3 kg
|  | Operating Environment |  | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation) | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Cyclone Injection Type Dry Measurement Unit
|  | Sample suction types | Cyclone type / One shot type / Hand shot type
|  | Suction nozzle | Can be selected from 3 types
Sampling Unit Specifications
|  | System | Cyclone type
|  | Communication method | USB (PC control)
|  | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Dimensions & weight | W 240×D 310×H 210 mm, 10 kg
|  | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)
Pressure Regulator Specifications
|  | Primary pressure | 0.6 to 0.8 MPa
|  | Secondary pressure | 0.05 to 0.5 MPa
|  | Filtration rating | Removing particles of 5 μm or larger
|  | Connection to air source | Tube of 6 mm outside diameter
|  | Communication method | USB (PC control)
|  | Required power supply | 115/230VAC(±10%), 100VA, 50/60Hz (excluding dust collector and compressor)
|  | Dimensions & weight | W 130×D 223×H 233 mm, 3 kg
|  | Operating Environment | Temperature: 10 to 30°C, Humidity: 20 to 80 % (no condensation)